## Sheet1
Company | Smartline |  |  |  |  | Cantel | Soluscope |  | AT-OS | MASS | Getinge | Steelco |  | Wassenburg | Torvan | YOMA | Notes | Darrens Notes and Queries?
Model | RotaScope iQ | SlidaScope iQ | RotaScope Classic | SlidaScope Classic | RotaScope CORE | ENDODRY | DSC8000 | AS300 | ENDO 3 | DAS-9P-HG | EDS8 | ED250 | ED200S | DRY300D | Scope Cabinet | Endo Drying Cabinet
Storage Method | Vertical | Horizontal Basket |  |  | Vertical | Horizontal Cassette | Horiztonal Tub | Vertical | Vertical | Vertical | Horizontal Basket | Vertical | Horiztonal Cassette | Vertical | Vertical | Horizontal Tray
Number scopes stored | 9 | 8 |  |  | 9 | 8 | 8 | 10 | 12 | 9 | 8 | 9 | 8 | 9 | 8 | 8 or 16
Manufactured in | Australia | Australia | Australia | Australia | Australia | Germany | France | France | Italy | USA | Sweden(?) | Italy | Italy | NL | Canada | China
EN16442 Standards
CE Registered
FDA listed | [16 empty] | Smartline pending distrubutor
7 Day storage approved
30 Day storage approved | [16 empty] | Many websites advertise 7 days however claim EN16442 which implies 30 days
Channel Monitoring
Foot print | 0.36m2 | 0.36m2 |  |  | 0.36m2 | 0.32m2 | 0.57m2 | 0.85m2 | 0.88m2 | 0.7m2 | 0.52m2 | 0.76m2 | 0.73m2 | 0.58m2 | 0.6m2 | 0.79m2
Pass through
Pass Through X-Ray | [16 empty] | Cardiac or any theatre that has lead lining usually has significant X-Ray use and would require X-ray shielding between the theatre and the reprocessing room
Can a standalone cabinet be converted to pass through or reconverted to standalone | [6 empty] | ? |  | ? |  |  | ? | ? | ? |  | ?
If Pass Through - does the pass through effectively stop airflow between rooms of high and low pressure | [5 empty] | ? | ? | ? | ? |  |  | ? | ? | ? | ? | ? | Often operating theatres and cleaning rooms have different pressures allowing rapid airflow to exchange between the rooms via the cabinet which is an infection control risk - some pass through units completely block air between rooms for best practice
Independent air supply | [16 empty] | Some regional Standards for Medical air supply states that high use devices including Drying Cabinets and power tools should be on a separate air supply – this is a big problem for many sites and especially for brownfield sites that need to have medical air supply installed
Electronic countdown | [16 empty] | Individual scope status allows the department to select oldest stored scope first reducing cost of reprocessing scopes
Can the data system allow remote access for support | [5 empty] | ? | ? | ? | ? |  | ? | ? | ? | ? |  | ? | Some cabinets allow the support company to remote log in either directly or via the hospital network – systems that allow remote login independently reduces problems with security and does not need special hospital IT department approvals
Can the data system allow for integration to other systems | [5 empty] | ? | Only to SoluScope IT DATA SYSTEM | Only to SoluScope IT DATA SYSTEM |  |  | T-DOC |  |  | ? |  |  | Some stored material may be better connected to an existing tracking or data management system – often matches the tracking system in CSSD – If an output can be enabled it can save the hospital running duplicate master systems for reporting
Does the inbuilt data system provide wash traceability including all AER brands? | [10 empty] | Via T-DOC | [5 empty] | Many cabinets only allow the scope traceability to be from the cabinet at loading point, which does not give information about the origin of the scope cleaning – the AER (Scope washer ID may provide traceability to the washer and user, date and time) Some systems have a data system that provides this traceability NOTE: Some cabinets may only get data feed from one cabinet - some system allow for data validation from all cabinets by identifying the load or unload from that cabinet via a local data device at the washer
Does the data system provide reporting on system status and errors | [16 empty] | Some data systems will log and print system errors in an easy to manage format that instantly identifies errors to the user via screen and or printed notification
Can the data system print cycle records for each scope processed at time of unloading | [16 empty] | When a scope is stored, some cabinets may produce a summary ticket or sticker that provides a summary of who washed the scope, which washer it came through, where it was stored and who unloaded the scope and usually states the "last use" time for that scope so that it can be assessed at bedside before use on a patient. The ticket or sticker may then be applied to a manual tracking system for traceability.
Can the provider remotely monitor the performance and system status to alert the hospital of breakdowns, failures or warnings | [5 empty] | ? | ? | ? | ? |  | ? | ? | ? | ? |  |  | Some cabinets data systems allow the service agent or supplier to remotely monitor the system for correct activity and to report issues via early warning or monitor if the system is needing extra maintenance eg Filters becoming blocked or air system flow rates out of parameters.
Scope transfer in baskets "Hands Free" | [16 empty] | Some washers can hold the scope in the same basket as storage which means that the basket is decontaminated along with the scope allowing for “No Touch" of the scopes after washing – scope connecting hoses may or may not be transferred also
Automated reminders for cleaning | [5 empty] | ? | ? | ? | ? |  | ? | ? | ? | ?
Automated reminders for servicing and PQ | [5 empty] | ? |  |  | ? |  | ? | ? | ? | ?
Cabinet LED light colour coding | [16 empty] | To alert of system faults, cleaning requirements, power outages and more
Customisable design to suit department needs | [16 empty] | Door hinge side, variable wall thickness pass through options